## Checks
Acceptance Criteria Number | Description
PAC-10 | patient age < 35
PAC-30 | breast implants present
SAC-30 | lmlo missing
SAC-40 | rmlo missing
SAC-50 | lcc missing
SAC-60 | rcc missing
SAC-70 | patient name is empty or inconsistent
SAC-80 | patient birth date is empty or inconsistent
SAC-81 | patient age inconsistent
SAC-90 | patient id is empty or inconsistent
SAC-100 | study instance uid is empty or inconsistent
SAC-110 | study date is empty or inconsistent
SAC-120 | files too large (>10 GB)
SAC-130 | no corresponding 2D for 3D file
SAC-131 | no corresponding 3D for 2D file
SAC-140 | manufacturer is empty or inconsistent
FAC-10 | dicom cannot be read
FAC-20 | sop class uid isn't 1.2.840.10008.5.1.4.1.1.1.2 or 1.2.840.10008.5.1.4.1.1.13.1.3
FAC-21 | modality not 'MG'
FAC-23 | body part examined not 'BREAST'
FAC-24 | lossy image compression
FAC-26 | quality control image
FAC-27 | burned in annotation
FAC-29 | image laterality not L or R
FAC-30 | wrong Patient Orientation values
FAC-80 | manufacturer missing
FAC-90 | manufacturer model name missing
FAC-100 | pixel data missing
FAC-101 | error in extracting pixel array
FAC-110 | presentation intent type is not For Presentation
FAC-130 | Image processing tags don't match manufacturer specifications
FAC-140 | manufacturer not supported
FAC-150 | file too large (>1.5GB)
FAC-160 | view not mlo or cc
FAC-170 | modified view
FAC-181 | partial view
FAC-200 | missing one of the following: SOP Instance UID Series Instance UID Study Instance UID Patient ID Study Date Patient Birth Date
FAC-210 | Transfer syntax must be one of the following: 1.2.840.10008.1.2.1 - Explicit VR Little Endian 1.2.840.10008.1.2 - Implicit VR Little Endian 1.2.840.10008.1.2.2 - Explicit VR BigEndian 1.2.840.10008.1.2.1.99 - Deflated Explicit VR Little Endian 1.2.840.10008.1.2.5 - RLE Lossless 1.2.840.10008.1.2.4.70 - JPEG Lossless (Process 14, SV1) 1.2.840.10008.1.2.4.80 - JPEG LS Lossless 1.2.840.10008.1.2.4.90 - JPEG2000 Lossless
IAC-20 | pixel intensity range insufficient
HOL-10 | image type
HOL-20 | estimated radiographic magnification factor
HOL-30 | photometric interpretation
HOL-40 | bits allocated
HOL-50 | bits stored
HOL-60 | high bit
HOL-80 | mismatch between WindowCenter WindowWidth and WindowCenterWidthExplanation
HOL-110 | grid
HOL-120 | focal spots
HOL-130 | exposure status
HOL-140 | rows
HOL-150 | columns
HOL-160 | window center
HOL-170 | window width
HOL-180 | pixel intensity relationship
HOL-190 | pixel intensity relationship sign
HOL-200 | presentation lut shape
HOL-210 | series description
GE-10 | image type
GE-20 | estimated radiographic magnification factor
GE-30 | photometric interpretation
GE-40 | bits allocated
GE-50 | bits stored
GE-60 | high bit
GE-70 | volumetric properties
GE-80 | window center and width explanation
GE-90 | voi lut function
GE-110 | grid
GE-120 | focal spots
GE-130 | exposure status
GE-140 | rows
GE-150 | columns
GE-160 | window center
GE-170 | window width
GE-180 | pixel intensity relationship
GE-190 | pixe intensity relationship sign
GE-200 | presentation lut shape